## Tabelle1
Specifications
No. | LOTs | Item | Specifications | Use | Quantity
[5 empty] | amount | unit
1 | Lot 1 | Automatic Laboratory Hydraulic Press | General Description: Preparation of pressed pellets for XRF, XRD, and other geochemical analyses. Compression of rock, mineral, and powder samples. See the illustrative image of a typical pressed pellet . Technical requirements : Compression Force : 40 t (400 kN) – 60 t (600 kN) Dwell Time: 0 to 999 seconds (adjustable) Maximum Piston Stroke: 50 mm Column Opening: requirement 220 mm Dimensions: Ergonomic design with a compact footprint. Power Supply: 220–400 V / 50–60 Hz / Alternating Current (AC) Power installed: 0.7- 1.4 kW Control Type: Fully automatic operation with a digital control panel Installation Requirements: Stable level floor; three-phase power supply; well-ventilated installation area Accessories / Consumables: Compression dies (standard diameters: 32 mm and 40 mm); replacement hydraulic oil, require certificates of conformity. User and Maintenance manuals in French. Verify compatibility with the electrical power supply (380 V three-phase available for hydraulic presses) and provide for any necessary adaptation works if required Reference Standards: Compliant with ISO and ASTM standards for geological sample preparation. Safety Standards: Compliant with applicable international and European safety standards (ISO / EN). | Pellet sample preparation for XRF, XRD, and other geochemical analytical techniques. Compaction pressure for rocks, minerals, and powders. | 2 | piece
|  | Final Destination | GIZ office Conakry/Guinea : GIZ- Programme Gouvernance des ressources du secteur extractif et des chaînes d’approvisionnement en Afrique de l’Ouest (GRSE) CORNICHE NORD X CITÉ MINISTERIELLE/DIXINN DI 256 - CONAKRY - GUINEA
2 | Lot 2 | Benchtop Energy Dispersive X-ray Fluorescence (EDXRF) Spectrometer | General description: Quantitative elemental analysis of rocks, minerals, soils, ores, cement, construction materials, and environmental samples. Elemental coverage from Fluorine (F) to Americium (Am). Technical Requirements: Advanced Configuration: Premium customized configuration equipped with advanced software options and a high-power excitation source, including: • 10 to 20 W / 40 to 60 kV / 2 to 6 mA X-ray tube • Dedicated anode (e.g., Silver (Ag) or Chromium (Cr)) • High-resolution Silicon Drift Detector (SDD) with enhanced Peltier cooling • Fundamental Parameters software • adequate module for multilayer analysis. Elemental Range: Fluorine (F) to Americium (Am), covering virtually all elements measurable by XRF Concentration Range: Sub-ppm to 100 wt% Sample Changer: Automatic removable sample changer with 8 to 15 sample positions Standard Sample Size: 32–40 mm diameter (approximately 2 inches) Detector: Silicon Drift Detector (SDD); resolution <135 eV; linear count rate >1,500,000 cps X-ray Tube: Silver (Ag) or Gold (Au) anode X-ray tube with automatic filter selection; power ≤15 W Measurement Atmosphere: Ambient air for most applications; helium optional for light-element analysis Environmental Compensation: Integrated pressure and temperature sensors for automatic correction Sample Spinner Included; improves analytical precision, particularly for heterogeneous samples, air filters, and particulate materials Software Instrument control, quantitative analysis using calibration standards, suitable standardless analysis, and FingerPrint identification module. Dedicated Computer Specifications: - Processor (CPU): A high clock frequency >3.6 GHz is recommended, as XRF spectral deconvolution algorithms are highly dependent on single-core performance. - Memory (RAM): 32 GB DDR5. Although 16 GB is the minimum requirement, 32 GB is recommended to ensure smooth processing of multiple comparative spectra simultaneously. - Storage: • Primary Drive (Operating System & Software): min 1TB (essential for rapid access to the software SQL database). • Secondary Drive (Backup & Archive): min 2 TB SSD or HDD for long-term storage of spectrum files (.spx) and audit trail records. Graphics Card (GPU): Mid-range dedicated graphics card (4 to 24 GB). A dedicated GPU reduces CPU workload during 3D visualization and spectral overlay rendering. Operating System: Microsoft Windows (64-bit). - User Interface: Dedicated external PC with an intuitive touchscreen interface. Power Supply: 100–240 V AC, 50–60 Hz (universal) Certifications: CE certified; IEC 61010-1 compliant (electrical safety); low X-ray radiation emission in compliance with applicable international regulations; require certificates of conformity. Key Geological Applications: Major and minor element analysis of rocks, bauxite, iron ores, limestone, fused glass beads, and pressed pellets User and Maintenance manuals in French. | Quantitative/qualitative elemental analysis of rocks, minerals, soils, ores, cements, construction materials, and environmental samples. Elemental coverage from fluorine (F) to americium (Am). | 1 | piece
|  | Installation, commissioning and training on-site | Destination: Conakry, Guinea (the End user laboratory) Onsite Training days: 1 day (8 hours) Number of participants: up to 10 participants Laboratory technicians and engineers Language of the training: French Goal of the training: Operator handover, safety, sample preparation, analytical procedures, calibration/quality, software use, basic maintenance, performance acceptance criteria. Include on-site training for technicians as a contractual condition, to be conducted at the LNG site in Conakry, Guinea. Other services Customer and maintenance service must be in French via email, telephone, or other virtual options.
|  | Final Destination | GIZ office Conakry/Guinea : GIZ- Programme Gouvernance des ressources du secteur extractif et des chaînes d’approvisionnement en Afrique de l’Ouest (GRSE) CORNICHE NORD X CITÉ MINISTERIELLE/DIXINN DI 256 - CONAKRY - GUINEA
3 | Lot 3 | Automatic Multi-Sample Thermogravimetric Analyzer (TGA) | General Description: Determination of mass loss as a function of temperature, including moisture content, ash content, volatile matter, fixed carbon, and Loss on Ignition (LOI). Suitable for the analysis of coal, coke, biomass, catalysts, soils, ores, cement, gypsum, and raw meal samples. Technical Requirements: Carousel Capacity: Automated simultaneous analysis of 15 to 20 samples End-of-Analysis Detection: Automatic, based on sample mass stabilization (constant weight criterion) Balance Accuracy: minimum 0.0001 g (0.1 mg) with a thermally isolated, low-drift analytical balance Furnace Atmosphere: Automated atmosphere control using air, nitrogen (N₂), or oxygen (O₂); programmable gas flow rates. Temperature Control: Programmable heating ramps with precise setpoint control and minimal temperature overshoot. Thermal Uniformity Rotating carousel to ensure uniform thermal exposure of all samples. Furnace Safety: Lockable furnace lid in the open position to prevent accidental closure in the event of power or compressed air failure Method Flexibility: User-configurable analytical methods capable of replicating conventional gravimetric procedures Software: Touchscreen interface with data management software and automatic report generation Power Supply: 220–240 V AC / 40 to 60 Hz Compliance Standards: ASTM D3172, ASTM D3173, ASTM D3174, ASTM D3175 (coal and coke analysis); ISO 11722, ISO 1171, and ISO 562 (or equivalent standards) Key Geological Applications: Loss on Ignition (LOI) determination for rocks and minerals; moisture and organic matter determination in soils; analysis of carbonate-bearing ores Require certificates of conformity. User and Maintenance manuals in French. | Analysis of mass loss of materials as a function of temperature: moisture, ash, volatile matter, fixed carbon, and loss on ignition (LOI). Applicable to coals, cokes, biomasses, catalysts, soils, ores, cements, gypsums, flours. | 2 | piece
|  | Installation, commissioning and training on-site | Destination: Conakry, Guinea (the End user laboratory) Onsite Training days: 1 day (8 hours) Number of participants: up to 10 participants Laboratory technicians and engineers Language of the training: French Goal of the training: Operator handover, safety, sample preparation, analytical procedures, calibration/quality, software use, basic maintenance, performance acceptance criteria. Include on-site training for technicians as a contractual condition, to be conducted at the LNG site in Conakry, Guinea. Other services Customer and maintenance service must be in French via email, telephone, or other virtual options.
|  | Final Destination | GIZ office Conakry/Guinea : GIZ- Programme Gouvernance des ressources du secteur extractif et des chaînes d’approvisionnement en Afrique de l’Ouest (GRSE) CORNICHE NORD X CITÉ MINISTERIELLE/DIXINN DI 256 - CONAKRY - GUINEA
4 | Lot 4 | Portable Analyzer (Lithium & Critical Elements) | General Description: Field and laboratory analysis of virtually all elements in the periodic table (H to U). Particularly suitable for the detection of lithium (Li), beryllium (Be), boron (B), carbon (C), fluorine (F), and sodium (Na), which cannot be measured using portable XRF analyzers. Designed for mineral exploration, geochemical investigations, forensic science, archaeology, and oil & gas exploration. Technical Requirements: Elemental Range: Hydrogen (H) to Uranium (U) – full periodic table coverage Spectral Range: 190–950 nm (triple spectrometer configuration) Spectrometers: Three integrated spectrometers providing extended wavelength coverage for the measurement of H, F, N, O, Br, Cl, Rb, Cs, and S. Lithium Detection: High-sensitivity lithium emission line at approximately 675 nm; detection limit: 2–5 ppm. Potassium (K) Detection: Dedicated potassium emission line free from iron interference, providing superior performance compared with conventional K emission lines. Laser Technology: pulse energy: 5–6 mJ; repetition rate: is the requirement 50 Hz; pulse width: 1 ns. Power Supply: Integrated rechargeable lithium-ion battery (internal or external charging system) and AC mains power adapter. Battery Life: More than 6 hours of continuous operation on a single charge. Primary Display: 3.5 to 4.3-inch high-brightness color touchscreen, readable in direct sunlight. Secondary Display: Rear display for rapid viewing of analytical results. Primary Camera: Integrated camera with laser target indicator for precise sample alignment before and during analysis. Macro Camera: Secondary camera for QR code and barcode scanning, photographic documentation, and report generation. Connectivity: Wi-Fi (IEEE 802.11b/g/n), Bluetooth (BR/EDR + BLE), integrated GPS, and USB Type-C. Operating System: operating system with an intuitive app-based user interface. Data Storage: SD card (32 GB minimum capacity). Processor: minimum 1.2 GHz; minimum 2 GB RAM; minimum 16 GB internal storage. Adequate Argon Purge System- Patented Purge technology with user-replaceable argon cartridge to improve the signal-to-noise ratio, particularly in the ultraviolet (UV) spectral region. Emission Window UV/Visible transparent quartz window with "Pyrex-strength" impact resistance. PC Software Desktop/tablet software for calibration editing, quantitative calibration curves, and alloy/grade library management. Microanalysis / Elemental Mapping Laser spot size: requirement 100 µm; point-by-point elemental mapping on a programmable measurement grid. Certifications: Compliant with applicable laser safety regulations; MSHA approved where applicable; calibration traceable to ISO/IEC 17025 requirements. Require certificates of conformity User and Maintenance manuals in French. | Field and laboratory analysis of all elements of the periodic table (H to U), including lithium (Li), beryllium (Be), boron (B), carbon (C), fluorine (F), and sodium (Na), which may require techniques other than portable XRF. Applicable to mineral exploration, geochemistry, forensics, archaeology, and oil exploration. | 1 | piece
|  | Installation, commissioning and training on-site | Destination: Conakry, Guinea (the End user laboratory) Onsite Training days: 2 days (16 hours) Number of participants: up to 10 participants Laboratory technicians and engineers Language of the training: French Goal of the training: Operator handover, safety, sample preparation, analytical procedures, calibration/quality, software use, basic maintenance, performance acceptance criteria. Include on-site training for technicians as a contractual condition, to be conducted at the LNG site in Conakry, Guinea. Other services Customer and maintenance service must be in French via email, telephone, or other virtual options.
|  | Final Destination | GIZ office Conakry/Guinea : GIZ- Programme Gouvernance des ressources du secteur extractif et des chaînes d’approvisionnement en Afrique de l’Ouest (GRSE) CORNICHE NORD X CITÉ MINISTERIELLE/DIXINN DI 256 - CONAKRY - GUINEA