## Requirements
GIZ Project “Supporting Moldovan communities in hosting Ukrainian refugees” in Moldova
Supply, delivery and installation of Medical Equipment for Mother and Child Centre and training of the operational staff
General implementation requirements:
A. Material and equipment requirements: 1) All material must be new, current in distribution (not a discontinued model if the successor model is already in distribution). All material must be fully inspected and tested to demonstrate compliance with specification requirements. (Manufacturer's Declaration of Conformity and Product Certificate). 2) The equipment must be supplied in accordance with the EC / DIN or ISO standards currently in force and manufactured in accordance with ISO safety rules and regulations. 3) The equipment shall be delivered in the original packaging and marked by the manufacturer in accordance with international requirements. 4) All equipment must comply with: • 230 V / 400 V 3-phase AC with grounding conductor. • 230 V AC with grounding conductor • 50 Hz frequency system 5) Recurrent unexpected interruptions of the power supply must not destroy the electrical and electronic circuits or mechanical parts of the instrument. 6) Measuring instruments, tables and dials as well as switches and hand tools must conform to the metric system. 7) Technical fluids and hydraulic oil (if any) must be suitable for moderate climates.
B. Documentation Requirements: 1) User manuals (Instructions for Use (IFUs), operator manuals), installation guides, risk management documentation and safety documentation shall be provided in English (Romanian or Russian if available) by the supplier/manufacturer, if applicable. 2) The manuals shall include, but not be limited to, all drawings for installation, general description of equipment, instructions for operation, maintenance, list of special tools, repair and adjustment, troubleshooting manuals, equipment lists, and spare parts lists, such as information necessary for ordering spare parts for all equipment furnished under the contract. 3) Labels and stickers of all types must be in English or Romanian with understandable symbols.
C. Quality Standards: 1) Products and services must comply with the international standards listed below or appropriate European standards. • IEC / ISO / EC (International Standard) • EN (DIN, VDE) (European Standard) 2) Unless otherwise specified herein, all materials used and equipment supplied and all workmanship and testing shall conform to the latest editions of ISO, IEC, EN, DIN and VDE standards. 3) Regulatory Approval / Certification must be submitted together with the offer 4) The equipment must meet all the legal conditions, certifications / permissions for operation on the local market.
D. Commissioning. 1) The equipment is handed over by the Supplier and received by GIZ and Beneficiary, when: a) The equipment is installed and the Beneficiary certifies that the system works without default and countersigns a Handover Act prepared by GIZ. b) The training is delivered and the Training Protocol is signed by all participants and Beneficiary. The training protocol shall be submitted to GIZ. c) The final Invoice and originals of documents are submitted to GIZ for payment.
E. Project location, climate and other information. Chișinău is the capital and largest city of the Republic of Moldova. Located along the Bîc River in the central region of the country, it serves as the primary economic, cultural, and administrative hub. Chișinău Municipality includes the city itself plus 34 surrounding suburban towns and villages. It spans about 572 km² with approximately 720,000 residents. Chișinău is the primary growth center in the country. Nearly one in three people in Moldova now reside within the Chișinău Municipality, marking a significant increase over the last decade as people move to the capital for economic and educational opportunities. The city experiences a humid continental climate (Köppen climate classification Dfb), characterized by four distinct seasons ranging from --10°C in winter to +30°C in summers. Annual precipitation is moderate, averaging roughly 500 to 550 mm (20 to 22 inches). Rainfall is fairly evenly distributed throughout the year, though summer often sees brief, heavy thunderstorms.

## Lot 1. USG
GIZ Project “Supporting Moldovan communities in hosting Ukrainian refugees”
Supply, delivery and installation of Medical Equipment for Mother and Child Centre and training of the operational staff
Lot 1. General Ultrasound System (1 item)
INDICATIVE TECHNICAL SPECIFICATIONS
DESCRIPTION: Diagnostic ultrasonography that uses ultrasound-based imaging technology for real-time visualization of anatomical structures, blood flow and functional parameters. The system is intended for multidisciplinary investigations, providing high-resolution images, advanced analysis functions and reliable results for routine examinations and complex cases.
1) Ports Sample | Minimum 4
Probe Type | Linear frequency band 4.8 - 16 MHz, active band or minimum field of view 25 mm
Convex frequency band 2 - 6 MHz, active band or minimum field of view 70°
Sectoral frequency band 2.3 - 6 MHz, active band or minimum field of view 90°
The device will be set to work with the supplied probes | Yes
2) Gray Levels
Preprocessing | digital channels ≥10 000 000 or 256 (physical channels)
Dynamic range summary | Minimum 350 dB
Maximum scanning depth | Minimum 35 cm
Possibility of image rotation | Minimum 180 degrees
Possibility of mirroring the image | Left / Right
POST-PROCESSING | Yes
3) Image modes | M-mode
2-D/B-Mode
3-D (automatic)
Harmonic imaging
Compressive elastography for the linear and convex probe with which the device will be delivered,
4). Doppler | Type PW, CFM, CW
Frequency display
Speed display
Power Doppler
HPRF yes
Real-time duplex yes
Real-time triplex yes
5). Capabilities | Digital measurements
Configuring measurements according to user needs
Selectable dynamic range
Continuous dynamic range focusing
Adjustable transmission focus
Measurements on video playback
Report or final protocol formation by the sonographer. With the possibility of change according to the needs of the end user.
Non-Doppler diagnostic mode for all probes with which the device will be delivered. | Yes
The presence of image enlargement technology without reducing the resolution, contrast and clarity of the area being enlarged or equivalent. | Yes
Automated image adjustment system for 2D/B-mode | Yes
Automatic adjustment of working frequencies by the device. | Yes
Trapezoidal imaging mode or widening of the scanning field for the linear probe | Yes
Automated measurements in pulsed Doppler (PW) mode | Yes
PAN/ZOOM real-time image | Yes
PAN/ZOOM frozen image | Yes
IMAGE STORAGE Capacity ≥ 1TB | Yes
Cine Memory Minimum 950MB or min 80000 frames or min 600 sec | Yes
DICOM 3.0 | Yes
USB input/output port | minimum 3 pcs
HDMI or DVI video output | Yes
LAN/NET RJ45 | Yes
Built-in battery with min. 2 h autonomy | Yes
6). Analysis packages | Transcranial applications
Abdomen
Soft parts
Vascular
Pediatrician
Musculoskeletal
Compressive elastography
7). Monitor | Minimum diagonal 23 inches
Minimum resolution 1920x1080
Articulated arm with height adjustment and right-left articulation
8). Monitor division | Mandatory presence of the control monitor
Touch screen technology
Diagonal Minimum 10 inches
9). Work console | Integrated digital keyboard / on console or sliding
Console lighting
Lighting intensity control
User programmable buttons
Work console Height adjustment
Presence of at least 4 wheels | Minimum 2 brakes
10). Accessories
Integrated black and white printer | Yes
Ultrasound gel | Minimum 1 L
Black and white printer paper roll | 5 rolls
Installation and Training: 1) The items shall be delivered and installed at the Public Health Institute of Mother and Child (IMC) address: 93 Burebista str. Chisinau, MD-2062. (IMSP Institutul Mamei și Copilului). 2) Installation, verification of equipment performance and start-up must be carried out by an authorised specialist certified by the manufacturer. 3) The training shall be carried out by the trainers authorized by the manufacturer or Supplier on general topics “Overview, operation, maintenance and troubleshooting” and “Periodic maintenance procedures and functionality / metrology testing procedures”. Staff: from 2 to 3 operators, bioengineers and technical specialists. Duration:1 day in total. Time will be coordinated with the supplier after installation. Venue: IMC premises. A Training Occurrence Protocol must be signed by all participants upon completion.

## Lot 2. Endoscopy
GIZ Project “Supporting Moldovan communities in hosting Ukrainian refugees”
Supply, delivery and installation of Medical Equipment for Mother and Child Centre and training of the operational staff
Lot 2. Rigid Endoscopy Image Processor (1 item)
INDICATIVE TECHNICAL SPECIFICATIONS
DESCRIPTION: The rigid endoscopy image processor is the central unit that receives the optical signal from the camera and converts it into high-resolution video images optimized for detailed visualization of anatomical structures. It integrates advanced algorithms for color correction and image noise reduction, providing maximum clarity during procedures.
A1) Technical parameters
Processed and displayed image quality | Minimum Full HD (≥1920 × 1080 pixels), preferably 4K UHD
Image acquisition technology | Minimum CMOS or superior
Video camera head | Included in the set; minimum IPX7 protection rating or equivalent
Optical adapter | Included in the set; compatible with rigid telescopes used in laparoscopic, gynecological and obstetric surgery
Optical compatibility | Compatible with standard rigid telescopes with C-Mount eyepiece or equivalent
Image brightness adjustment | Automatic Exposure control and automatic image brightness optimization
Image enhancement | Digital image optimization functions (detail enhancement, noise reduction, color adjustment)
White balance | Automatic and/or manual
Zoom | Digital and/or optical
Freeze | Image freeze function
Display type | LCD and/or display of information on the medical monitor
Photo and video storage | USB and/or internal memory
Auto focus | Yes
Telescope connection | Compatible with standard rigid telescopes used in minimally invasive surgery
Camera buttons | Programmable – minimum 2
Camera cable length | Minimum 2.5 m
Disinfection / sterilization | According to manufacturer's recommendations
Video outputs | Minimum HDMI; DVI-D, DisplayPort or equivalent accepted
Resolution compatibility | Compatible with Full HD and/or 4K monitors
A2). Light Source
Light source technology | LED or superior
Light source lifetime | Minimum 20,000 hours
Light intensity adjustment | Yes
Light guide / optical fiber | Minimum 1 unit; preferably a second spare light guide
Light source compatibility | Compatible with the offered video system and telescope
A3). Medical Monitor
Medical monitor size | Minimum 24 inches, preferably 27 inches or larger
Medical monitor resolution | Minimum Full HD, preferably 4K
Medical-grade monitor | Yes
Monitor compatibility | Compatible with the offered video system and telescope
A4). Medical trolley/cart
Medical trolley/cart | Capable of accommodating all requested modules
Installation: 1) The items shall be delivered and installed at the Public Health Institute of Mother and Child (IMC) address: 93 Burebista str. Chisinau, MD-2062. (IMSP Institutul Mamei și Copilului). 2) Installation, verification of equipment performance and start-up must be carried out by an authorised specialist certified by the manufacturer. Training is not nececessary for this Lot.

## Lot 3. Colonoscope
GIZ Project “Supporting Moldovan communities in hosting Ukrainian refugees”
Supply, delivery and installation of Medical Equipment for Mother and Child Centre and training of the operational staff
Lot 3. Video Colonoscope (1 item)
INDICATIVE TECHNICAL SPECIFICATIONS
DESCRIPTION:Flexible videocolonoscope for detailed visualization of the mucosa during videocolonoscopy procedures for adults and children aged 10 to 18 years.
Special requirement: | Fully compatible with the existing Olympus CV-1500 video processor without the need for additional adapters.
Technical parameters
Patient category | Adults and children aged 10 to 18 years
Working length | Minimum 1680 mm
Insertion tube diameter | Maximum 12.8 mm
Instrument channel diameter | Minimum 3.7 mm
Working channel with suction function | Yes
Auxiliary lens and/or operative field irrigation function | Yes
Up/Down angulation | Minimum 180 degrees
Right/Left angulation | Minimum 160 degrees
Field of view | Minimum 140 degrees
Depth of field | 2–100 mm
Displayed image quality | Minimum HDTV
Special functionality | Compatible with Narrow Band Imaging (NBI or equivalent) technology provided by the video processor
Installation: 1) The items shall be delivered and installed at the Public Health Institute of Mother and Child (IMC) address: 93 Burebista str. Chisinau, MD-2062. (IMSP Institutul Mamei și Copilului). 2) Installation, verification of equipment performance and start-up must be carried out by an authorised specialist certified by the manufacturer. Training is not nececessary for this Lot.

## Lot 4. Anesthesia Machine
GIZ Project “Supporting Moldovan communities in hosting Ukrainian refugees”
Supply, delivery and installation of Medical Equipment for Mother and Child Centre and training of the operational staff
Lot 4. Anesthesia machine (advanced features) (1 item)
Description | The anesthesia machine is designed to deliver and monitor anesthetic gases and provide artificial respiration for the patient during surgery
Parameter |  |  | Specification
Gas outlets |  |  | O2, Air
Anesthesia machine display |  |  | ≥15", color TFT or LCD
touch screen
Flow meters | type |  | electronic
| gas |  | O2, Air
| range, L/min |  | range min. 0.2 - 15
Vaporizer | Vaporizer type accepted | Isoflurane | required
|  | Sevoflurane | required
|  | Halothane | required
| Number of vaporizers installed on the device | ≥ 2 units | required
|  | Isoflurane | Required
|  | Sevoflurane | required
| Interlock |  | required
| Absorption system |  | required
Safety mechanisms | O2 safety |  | acoustic, visual
| hypoxic mixture safety |  | Required
Automatic ventilator | Patient type |  | Adult, Pediatric, Neonatal
| Ventilation modes |  | Manual/spontaneous, VCV, PCV, PSV/PS, SIMV-V, SIMV-P
| Electronic gas mixing mechanism (mixer) |  | Required
| Tidal volume, ml |  | range min. 5–1500
| Respiratory rate/minute |  | range min. 5–100
| Minute volume (MV), L/min |  | range min. 3–60
| I:E ratio |  | minimum 2:1 to 1:8
| Inspiration pause |  | required
| pressure limit, cmH2O |  | range min. 10–70
| PEEP, cmH2O |  | range min. 4–30
| Flow trigger |  | range min. 1–10 L/min
| Pressure trigger |  | range min. 20–1 cmH₂O
| Compliance and resistivity monitoring |  | required
| Self-diagnostic system |  | leak testing, respiratory circuit testing, compliance, gas supply, verification of all systems
AGSS (anesthetic gas scavenging system) |  |  | mandatory
Anesthetic agent consumption management system |  |  | optional
Pneumatic patient ventilation circuit | with gas mixture heating function |  | Required
| Auxiliary gas mixture outlet port |  | Required
Parameters monitored and displayed on the screen | Air pressure | High pressure alarm | required
|  | Low-pressure alarm | Required
|  | Low pressure / apnea | Required
|  | Other pressure alarms | Required
| Expiratory volume / flow |  | required
| Minute volume, l/min |  | required
| O2 concentration | Response time, sec | ≤ 35
| CO2 | apnea alarm | required
| Agent monitoring | Type of agents | Isoflurane, sevoflurane, halothane (optional)
|  | Automatic identification of anesthetic gases | Required
|  | Agent concentration alarm | Required
|  | Determination and display of MAC value | Required
| Spirometry |  | optional
Gas module | built into the anesthesia machine |  | required
| measures gas concentrations: O2, CO2, anesthetic agents |  | required
| Paramagnetic O2 sensor |  | required
Vital signs monitor | Monitor | ≥19", color TFT or LCD | Required
|  | touch screen | required
| monitor dedicated to displaying vital signs |  | required
| mounting the monitor to/from the anesthesia machine |  | required
| thermal printer |  | required
| WiFi/LAN/HDMI/USB connections |  | Required
| 360J defibrillation protection |  | required
| Protection against high-frequency electrosurgical devices |  | Required
| Automatic isoline drift compensation |  | Required
| Automatic cuff type detection |  | Required
| Internal rechargeable battery |  | required
| Communication interface with other devices |  | required
Included hemodynamic modules | Electrocardiogram (ECG) | Heart rate | required
|  | ECG trace | required
|  | ST segment analysis and measurement | required
|  | Identification of at least 20 arrhythmias | required
| Pulse oximetry (SpO2) | photoplethysmography | required
|  | SpO2 value | required
|  | perfusion index | required
| Non-invasive blood pressure (NIBP) |  | required
| Respiration (transthoracic impedance) |  | required
| 2-channel temperature |  | required
| Invasive blood pressure (IBP) on 2 channels |  | required
| BIS (bispectral index) monitoring mode or objective monitoring mode for the depth of intra-anesthetic neuromuscular block (TOF/NMT) |  | required
Priority alarms |  |  | 3
Power supply voltage |  |  | 220 V, 50 Hz
Auxiliary 220 V outlets ≥ 3 |  |  | required
Internal battery | Rechargeable |  | required
| Operating time ≥ 1.5 hours |  | required
Storage drawer ≥ 2 pcs |  |  | required
Centralized brake for securing the device or classic locking system for at least two wheels |  |  | required
Gas supply pressure |  |  | 3.0 - 6 bar
Accessories
Hose with connector for connection to the compressed air source |  |  | 1 pc.
Hose with connector for connection to the oxygen source |  |  | 1 pc.
Ventilation circuit | Adult, reusable |  | ≥ 2 sets, specify product code
Ventilation circuit | Neonatal pediatric, reusable |  | ≥ 2 sets, specify product code
Pediatric test lung | pediatric, reusable |  | ≥ 2 pcs., indicate product code
Flow sensor | Reusable |  | ≥ 2 pcs., indicate product code
Antibacterial filter | Adult, single-use |  | ≥ 200 pcs., please specify product code
Gas module accessories | Adult |  | ≥ 2 sets, specify product code
ECG cable |  |  | ≥ 2 pcs., specify product code
ECG electrode | Adult, single-use |  | ≥ 100 pcs., indicate product code
ECG electrode | Pediatric, single-use |  | ≥ 100 pcs., indicate product code
SpO2 sensor | Adult, reusable |  | ≥ 2 pcs., indicate product code
SpO2 sensor | Pediatric, reusable |  | ≥ 2 pcs., specify product code
SpO2 sensor | Newborn, reusable |  | ≥ 2 pcs., please specify the product code
NIBP cuffs | Adult, reusable |  | ≥ 2 pcs., indicate product code
| Large adult, reusable |  | ≥ 2 pcs., indicate product code
| Large pediatric, reusable |  | ≥ 2 pcs., indicate product code
| Pediatric, reusable |  | ≥ 2 pcs., specify product code
| Small pediatric, reusable |  | ≥ 2 pcs., indicate product code
Temperature sensor | Adult, reusable |  | ≥ 2 pcs., indicate product code
Temperature sensor | Pediatric, reusable |  | ≥ 2 pcs., specify product code
Accessories required for operation of the BIS or TOF/NMT module | Adult accessories |  | ≥ 5 pcs., specify product code
Installation and Training: 1) The items shall be delivered and installed at the Public Health Institute of Mother and Child (IMC) address: 93 Burebista str. Chisinau, MD-2062. (IMSP Institutul Mamei și Copilului). 2) Installation, verification of equipment performance and start-up must be carried out by an authorised specialist certified by the manufacturer. 3) The training shall be carried out by the trainers authorized by the manufacturer or Supplier on general topics “Periodic maintenance and disinfection procedures and functionality / calibration and testing procedures”. Staff: from 3 to 4 operators, bioengineers and technical specialists. Duration:1 day in total. Time will be coordinated with the supplier after installation. Venue: IMC premises. A Training Occurrence Protocol must be signed by all participants upon completion.