## Specifications
Tender document
Specification | [5 empty] | Entry by the bidder
Item No | Specification | Unit | Quantity | Installation and Commissioning (Yes/No) | Training (Yes / No) | (Please confirm the specification with "Yes" and tell us the brand name, model, country of origin and other information about the offered product)
1 | Microbial identification system | item | 1
1.1 | General requirements for the product: System for rapidly identifying microorganisms by creating a "mass spectrum" fingerprint from a sample and comparing it to a database of known microbial spectra. The process involves mixing a small amount of a microorganism, often from a single colony, with a matrix solution and a laser. The laser desorbs and ionizes the sample's proteins, and the ions are then separated by their mass-to-charge ratio (\(m/z\)) in a time-of-flight tube. The resulting spectrum is compared to a database to identify the organism, with applications in clinical, food, and environmental microbiology.
1.2 | Technical data: Dimensions • W*D*H approximately: 70*50*110 cm • Benchtop solution: weight approx. 75 kg Performance • Mass range not less than 200,000 Da • Proteins mass range resolution - Insulin (m/z 5,734 Da) ≥ 400 FWHM - Myoglobin M2+ (m/z 8,476 Da) ≥ 600 FWHM - Cytochrome C (m/z 12,361 Da) ≥ 700 FWHM - Myoglobin (m/z 16,952 Da) ≥ 800 FWHM • Mass accuracy (measured on protein mixture) - with external calibrant better than 200 ppm - with internal calibrant better than 150 ppm • MS sensitivity BSA (m/z 66,000 Da) 500 fmol at S/N ≥ 50:1 shown with 500 laser shots General description: Fully automated benchtop linear mass spectrometer with a target plate capacity of not less than 45 sample spots in a single run. • Ion acceleration: Positive ion mode up to +20 kV • Effective flight path not less than 95 cm • Max. data acquisition speed of 200 Hz sufficient for routine microbiological diagnostics • Laser system • solid-state laser - wavelength 355 nm - repetition rate not less than 200 Hz - lifetime minimum 6 years or not less than 450 million shots, whatever comes first • Vacuum system - vacuum pump turbomolecular
1.3 | (Standard) Accessories: Consumables / Reagents available • Small assay volume and homogenous spot generation for MALDI • sample clean up enables the reduction of salt interferences • Active Drying for homogenous spot shapes • accessories supply for sample preparation • Modular and customized automation solutions • Reagents to prepare samples for analysis, ensuring the analyte can be ionized and detected by the mass spectrometer. • Sample preparation kits • Kits for specific analyses, such as lipid, peptide, or protein imaging • Kits for pre-treating specific organisms like blood cultures or filamentous fungi • Target plates • Calibration standards • Kits enable easy and convenient preparation of matrix solutions for use in MALDI Imaging • Consumables for at least 1000 tests included. Spare Parts • Matrices, • Targets, • Sample preparation kits, • Internal components, • Reflectron and detector. |  | standard
1.4 | Installation: - Installation, verification of equipment performance and start-up in the laboratory must be carried out by an authorised specialist certified by the manufacturer. - After installation and start-up, the supplier must check the analytical performance of the system and certificate that the system works without default. -The equipment must be supplied with the consumables and spare parts required for installation, commissioning and provisional acceptance, as well as for the training and assistance provided. |  |  | yes
1.5 | Training details/conditions: The training covers system overview, calibration, reagent/sample preparation, test execution, data analysis, compliance, maintenance, and troubleshooting. Hands-on sessions on sample loading, protocol selection, mass spectrometry principles, quality control and LIS integration. Users will also learn contamination prevention measures, troubleshooting common errors. Target Staff: • Laboratory Technologists – Operation, test execution & troubleshooting • Clinical biochemists – Data analysis & validation • Quality Assurance (QA) Staff – Compliance, documentation & quality control • Biomedical Engineers – System maintenance, calibration & error handling Number of trainees: About 5 - 6 • Estimated time: 8 hours • Location: Ukrainian Laboratory in Dnipro. • Language: Ukrainian or English with translation into Ukrainian (provided by the supplier) • A Training Occurrence Protocol must be signed by all participants upon completion | Piece | 1 |  | yes
1.6 | Transport:DPU Dnipro, Ukraine
General Information
Project location, climate and other information: Dnipro, Ukraine. Operating conditions at planned location • enviromental conditions: 16°C to 30°C; 20% to 75% non-condensing • Operating noise level ≤ 65 dB
Mandatory requirements for all material items
Standards Documentation Requirements in English and in Ukrainian Language. : •User manuals (Instructions for Use (IFUs),operator manuals), installation guides, risk management documentation and safety documentation •equipment lists and spare parts lists •Applicable calibration certificates. Regulatory Approval / Certification : Must be submitted together with the offer. • CE Marking or equivalent •the supplier shall hold a valid ISO 13485 certificate or equivalent • instrument will be used for patient diagnostics (IVD certified) - згідно технічного технічного регламенту МОЗ України івд розділяють на 4 категорії за рівнем ризику (дод 2). Тести крові на віл снід, тести на вагітність потребують цього, достатньо самодекларування виробника. International Standards • IEC 61010-1:2020 Safety requirements for electrical equipment for measurement, control, and laboratory use • IEC 61010-2 Safety requirements for electrical equipment for measurement, control, and laboratory use • IEC 61326-1 EMC • IEC 61326-2 EMC • Traceable to NGSP and IFCC reference methods • CLSI LIS1-A and LIS2-A compliant interface • ISO 10927:2018 standard for polymer analysis
Material and equipment requirements All material must be new, current in distribution (not a discontinued model if the successor model is already in distribution). All material must be fully inspected and tested to demonstrate compliance with specification requirements. (Manufacturer's Declaration of Conformity and Product Certificate). Connectivity • Network interface, minimum RS232-Interface port • Flexible LIS connectivity. Electrical Supply • Electrical supply 100 - 240V AC, 50Hz Software/Hardware Upgrade Availability • Periodic software updates for performance optimization. • Future hardware upgrades for throughput or connectivity improvements
Aftersales services: Repair and maintenance services, as well as spare parts, for the product offered, are available in the country of use. This includes (in particular) - Technical support like Hotline / email support,Remote diagnostics (often via secure connection), Troubleshooting assistance,Software updates & patches - Maintenance services of the delivered product in accordance with the manufacturer's specifications. It is assumed that at least one maintenance service will be required within the first year of use. - Repair services of the delivered product in accordance with the manufacturer's specifications - Provision of OEM spare parts for the delivered product Bidders are required to expressly confirm compliance with this requirement in their bid (PDF) and provide the address(es) of the service center(s) in Ukraine.