## Lot 1
No. | Item | Specification | Quantity
1 | Meter Assembly | To incorporate a removable strainer at the inlet. Strainer close fitting and designed for easy removal. Inlet strainers shall be of nickel-plated copper, stainless steel or other materials having satisfactory characteristics | 2250
To incorporate a non-return valve within the meter
2 | Meter Box | Meter box to contain Quarter turn Valve
Should be UV resistant
Should be secure and lockable
Should have drainage holes to prevent ponding or flooding or moisture build up.
3 | Meter Body | The meter body shall have a meter counter seal (anti-fraud seal)
Made from Composite, UV stabilized Copolymer.
The year of manufacture shall be indicated on the body
The meter body shall have no external calibration screw to prevent fraud.
Both the inlet and outlet shall be the same size and on the same axis. The corresponding fixed set of dimensions to BS 5728/1, ISO 4064 – 1: 2024, or equivalent
The water meter shall be threaded. The meters shall be supplied complete with a set of connectors that are made of copper alloy or equivalent material resistant to corrosion, rust and damage due to shock or vibration. The connectors shall be threaded to the correct male size, comprising cap nuts, linings and fiber sealing washers. The meter linings shall have adequate provisions to safeguard against tampering.
4 | Marks and inscriptions | Manufacture’s Trademark
Year of Manufacture
Direction of Flow by Means of an Arrow on the body
MID water meter approval number
Printed Serial No.
Barcode or QR Code
Nominal Flow rate Q3
Class ratio Q3/Q1
Nominal Diameter
Temperature Class
Maximum Admissible Pressure (MAP)
Shall be engraved the name of the water utility company e.g. SWSC, EWSC, ChWSC, NWSC, LgWSC
5 | Meter – Hydraulic Chamber | Measurement Principal: Multi – Jet
Installation: The meter shall be approved for horizontal
Maximum admissible Pressure is 16 bar.
6 | Meter – Register | The meter register shall be extra -dry dial register with magnetically shielded against fraudulent tempering.The register is fully – sealed, vacuum filled, with a sample straight – reading presentation. The register is placed in a window made of glass with internal sealing gasket to prevent the formation of condensation inside the register. It must be installed in the meter body in the direction of flow for easy reading.
The register must have minimum of 5 dials (colour black) for cubic meters, as well as 2 pointers (colour red) for litres.
The register shall be pre-equipped for AMR.
The register shall be IP68 and the certificate shall be attached to the offer.
7 | Metrology | The meter size shall be DN15 and length of approx. 165mm.
The nominal flow rate shall be Q3 (m3/h) = 2.5
Class R160 MID approval certificate shall be attached to the offer.
The Maximum Permissible Error (MPE) for the upper flow rate zone (Q2 ≤ Q ≤ Q4)
shall be ±2 %, (for temperatures from 0.1 °C to 50 °C and ±3 % for Temperatures greater than 50 °C).
The MPE for the lower flow rate zone (Q1 ≤ Q <Q2) shall be ±5 %
Regardless of the temperature range.
Maximum allowable hydraulic pressure loss across the meter to be ΔP = 0.063 MPa
The meter to conform to T30 temperature range in accordance with MID standards
Meter body should allow for minimal serviceability such as changing strainers or Non–return valve (nrv) where applicable.
8 | Certification | Meters shall comply with ISO 4064:2024 and/or OIML R49-1:2024 and MID certificate Directive 2014/32/EU, or equivalent (Provide proof) All materials in contact with potable water shall have valid sanitary (hygienic) certification such as WRAS, NSF/ANSI 61, KTW, or equivalent. Manufacturer’s Authorization to sell their products issued by the original equipment manufacturer is required. Factory verification/test certificates (batch or individual) are required. Offered products must be manufactured under ISO 9001:2015. Manufacturer of the offered products must be ISO 14001:2105 certified. Manufacturer of the offered products must be ISO 45001 certified.

## Lot 2
No. | Item | Specification | Quantity
1 | Meter Assembly | To incorporate a removable strainer at the inlet. Strainer close fitting and designed for easy removal. Inlet strainers shall be of nickel-plated copper, stainless steel or other materials having satisfactory characteristics | 2250
To incorporate a non-return valve within the meter
2 | Meter Box | Meter box to contain Quarter turn Valve
Should be UV resistant
Should be secure and lockable
Should have drainage holes to prevent ponding or flooding or moisture build up.
3 | Meter Body | The meter body shall have a meter counter seal (anti-fraud seal)
Made from Composite, UV stabilized Copolymer.
The year of manufacture shall be indicated on the body
The meter body shall have no external calibration screw to prevent fraud.
Both the inlet and outlet shall be the same size and on the same axis. The corresponding fixed set of dimensions to BS 5728/1, ISO 4064 – 1: 2024, or equivalent
The water meter shall be threaded. The meters shall be supplied complete with a set of connectors that are made of copper alloy or equivalent material resistant to corrosion, rust and damage due to shock or vibration. The connectors shall be threaded to the correct male size, comprising cap nuts, linings and fiber sealing washers. The meter linings shall have adequate provisions to safeguard against tampering.
4 | Marks and inscriptions | Manufacture’s Trademark
Year of Manufacture
Direction of Flow by Means of an Arrow on the body
MID water meter approval number
Printed Serial No.
Barcode or QR Code
Nominal Flow rate Q3
Class ratio Q3/Q1
Nominal Diameter
Temperature Class
Maximum Admissible Pressure (MAP)
Shall be engraved the name of the water utility company e.g. SWSC, EWSC, ChWSC, NWSC, LgWSC
5 | Meter – Hydraulic Chamber | Measurement Principal: Multi – Jet
Installation: The meter shall be approved for horizontal
Maximum admissible Pressure is 16 bar.
6 | Meter – Register | The meter register shall be extra -dry dial register with magnetically shielded against fraudulent tempering.The register is fully – sealed, vacuum filled, with a sample straight – reading presentation. The register is placed in a window made of glass with internal sealing gasket to prevent the formation of condensation inside the register. It must be installed in the meter body in the direction of flow for easy reading.
The register must have minimum of 5 dials (colour black) for cubic meters, as well as 2 pointers (colour red) for litres.
The register shall be pre-equipped for AMR.
The register shall be IP68 and the certificate shall be attached to the offer.
7 | Metrology | The meter size shall be DN15 and length of approx. 165mm.
The nominal flow rate shall be Q3 (m3/h) = 2.5
Class R160 MID approval certificate shall be attached to the offer.
The Maximum Permissible Error (MPE) for the upper flow rate zone (Q2 ≤ Q ≤ Q4)
shall be ±2 %, (for temperatures from 0.1 °C to 50 °C and ±3 % for Temperatures greater than 50 °C).
The MPE for the lower flow rate zone (Q1 ≤ Q <Q2) shall be ±5 %
Regardless of the temperature range.
Maximum allowable hydraulic pressure loss across the meter to be ΔP = 0.063 MPa
The meter to conform to T30 temperature range in accordance with MID standards
Meter body should allow for minimal serviceability such as changing strainers or Non–return valve (nrv) where applicable.
8 | Certification | Meters shall comply with ISO 4064:2024 and/or OIML R49-1:2024 and MID certificate Directive 2014/32/EU, or equivalent (Provide proof) All materials in contact with potable water shall have valid sanitary (hygienic) certification such as WRAS, NSF/ANSI 61, KTW, or equivalent. Manufacturer’s Authorization to sell their products issued by the original equipment manufacturer is required. Factory verification/test certificates (batch or individual) are required. Offered products must be manufactured under ISO 9001:2015. Manufacturer of the offered products must be ISO 14001:2105 certified. Manufacturer of the offered products must be ISO 45001 certified.